Testing the High-Power Switch Board

Servicing a SuperDARN Transceiver
Step 8

Content

1.	Introduction	3

2.	Instructions	3

3.	Shortcut	5

4.	Conclusion	6

- Introduction

This document provides work instructions for testing the High-Power Switch Board in a SuperDARN transceiver box. Before attempting to implement these instructions, be sure to complete all the preceding steps in the procedure for Servicing a SuperDARN Transceiver.

- Instructions

Following are the step-by-step instructions for testing the High-Power Switch Board. In the case of unforeseen problems occurring, apply electronic fault-finding techniques. Refer to Figure 1 for component and connection locations.

- Switch on the AWG. Set the frequency to 12.5 MHz and the amplitude to 500 mVpp. Connect a BNC cable between the output of the AWG and Channel 1 on the Agilent MSO6104A to view the waveform. Press the Output button on the front face to enable the waveform. Remember to set the impedance of the channel to 50 Ω. Check that the waveform displayed is correct.

- Connect the Special Test Cable (STC) to the “Tx/Rx” of the Radar Lab 2.0 via the BNC connector and insert the other end into P2 on the Power Distribution Board. This will give us manual control over some of the FPGA functions needed for the tuning process.

- Terminate the output at the antenna port with a 100W dummy load.

- Make sure that the Phoenix connector J2 is removed from the HPSW. Replace it with a connection to the bench PSU. Set the PSU to 30 V and limit the current to 1 A.

- Switch on the power of the transceiver box, the Radar Lab 2.0; as well as the 30 V from the PSU to the HPSW.

- Check that the Tx/Rx pulses are coming through on P5.

- Check that TP_A is a replica of the Tx/Rx signal – see Figure 1. If there is no waveform on TP_A, check that the CPLD is programmed. Otherwise, there is possibly a fault with the CPLD or its clock.

Figure 1. HPSW Oscilloscope measurements for step 6.

- Check that the P6 switch waveform goes from a DC offset of about half of the 15 V supply (7.5 V) up to the HV supply voltage, which in this instance is set to 30 V.

Figure 2. HPSW Oscilloscope measurements for step 8.

- Power off.

- Replace the 30 V with the 1000 V (reconnect the phoenix connector from the HV supply board). Switch on and check that P6 is now the switched 1000 V.

- Power off and return to 30 V supply.

- Remove the dummy load and connect the RF signal to the antenna port with a frequency of 12.5 MHz and an amplitude of 0 dBm.

- Make sure that you terminate the input port from the power amp or leave it connected to the power amp.

- Power on.

- Check that the signal from the antenna can be seen at the receive port when the Tx/Rx switches low. You can also check that there is no leakage into the input port by terminating the receiver port and measuring on the input.

Figure 3. HPSW Oscilloscope measurements for step 15.

- Power off.

- Terminate the receiver port and connect the RF signal to the input.

- Power on.

- Measure that the signal is getting through to the antenna port when the Tx/Rx signal is high. Also check that there is no leakage to the receiver by terminating the antenna port and measuring on the receiver port.

Figure 4. HPSW Oscilloscope measurements for step 19.

- Shortcut

Once you are comfortable with what needs to be terminated when, where and why you can also use this shortcut testing sequence. If any problems occur, revert to the full testing procedure.

- Connect the HPSW (J1) to the PDB.

- Make sure that the Phoenix connector at J2 (HV Supply) is removed from the HPSW. Replace it with a connection to the bench PSU. Set the PSU to 30 V and limit the current to 1 A.

- Connect the Rx and PWRAMP RF connectors each to a channel on the oscilloscope (channels 2 and 3, set to a 50 Ω impedance).

- Connect the RF signal from the signal generator to the antenna port, but do not switch it on yet (12.5 MHz frequency and 0 dBm amplitude).

- Turn on the mains, 15 V and 30 V power.

- Check that Rx/Tx signal is coming through to the HPSW on pin P5 and P7 (TP_A). If you do not get a Tx/Rx signal here it is likely that the CPLD is either not programmed or broken.

- USE A HIGH VOLTAGE PROBE to check on P6 (it’s not HV yet, but it’s good practice) that your voltage goes from around 7 V to 30 V as Tx/Rx triggers.

- Switch on the AWG and set the frequency to 12.5 MHz and set the amplitude to 0 dBm. Connect a BNC cable between the Output of the AWG and Channel 2 on the Agilent MSO6104A to view the waveform (±640 Vpp).

- The settings used on the pre-set function generator can be selected instead of setting them.

- Connect the Special Test Cable (STC) to the Tx/Rx of the Radar Lab 2.0 via the BNC-connector and insert the other end into P2 of the Power Distribution Board. This will give us manual control over some of the FPGA functions needed for the tuning process. (This should already be done after completing the previous steps).

- Now switch on the signal generator output.

- Monitor the Rx/Tx signals on channels 2 and 3 on the oscilloscope.

  - Make sure that both the channels are set to a 200 mV scale on the oscilloscope, and they should have the same pk-pk voltages.

  - If Rx is more than 75 mV less than Tx; the problem is most probably with a faulty limiter (D4) for version 1.1B board.

  - Test it and, if necessary, replace it.

  - Other known culprits to be investigated (if it is found not to be the limiter) are the Q3 and Q4 MOSFETs, located on the bottom of the HPSW board.

Figure 5. HPSW Oscilloscope measurements for steps 11 and 12.
Red: 30 V, Yellow: TP_A
Green: Rx, Purple: Tx.

NOTE:
With the new HPSW design (1.1C) there are stronger limiter diodes in place, and they should not be the cause, but check anyway.

- Switch off the output of the signal generator, the 30 V supply, the front panel, and the mains.

- Switch off the output of the signal generator, the 30V supply, front panel, and mains.

- Remove the RF signal and connect the dummy load to the antenna port.

- Remove the oscilloscope channel connectors from the Rx and PWRAMP RF connectors and reconnect all the normal RF circuit connections.

- Replace the 30 V supply with the 1000 V (reconnect the phoenix connector from the HV supply board).

- YOU HAVE REMOVED THE FUNCTION GENERATOR FROM THE ANTENNA PORT, RIGHT!?

  - If not, you are about to pump around 1000 V into it. Make sure the dummy load is connected and NOT the function generator!

  - Also, make sure that the oscilloscope is disconnected from the Tx/Rx signals!

- Switch on the front panel, then switch on the 3.3 V supply connected with the STC.

- Use a HIGH VOLTAGE PROBE to check that P6 is now switched to around 850 V.

- Switch off the 3.3 V signal, then the front panel and the mains.

- Conclusion

This concludes the work instructions for testing the High-Power Switch Board of a SuperDARN transceiver box. The next step in the procedure for Servicing a SuperDARN Transceiver is to test the Power Amplification Board.